## Specifications
Product | Corival cpet
Partnumber | 960900
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Workload range fixed torque | Workload range in fixed torque mode | 0,1 - 70 Nm |  | x
| Minimum load | The minimum load the ergometer can provide | 7 W |  | x
| Maximum peak load | The maximum load the ergometer can provide for a short period of time | 1000 W |  | x
| Minimum load increments | The smallest steps with which the load can be added | 1 W |  | x
| Maximum continuous load | The maximum power the ergometer can deliver continuous | 750 W |  | x
| Hyperbolic workload control | Control of workload in hyperbolic way | yes |  | x
| Linear workload control | Control ergometer in a linear way | yes |  | x
| Fixed torque workload control | Control ergometer in a fixed torque way | yes |  | x
| Maximum rpm independent constant load | Rpm as to which constant load can be applied | 150 rpm |  | x
| Minimum rpm independent constant load | Rpm as from which constant load can be applied | 30 rpm |  | x
| Optional heart rate controlled workload | Control the device optionally through a set heartrate | yes |  | x
| Electromagnetic "eddy current" braking system | System that very accurately applies a brake to the ergometer | yes |  | x
| Dynamic calibration | Continuous calibration during application of workload | yes |  | x
| Power range at maximum rpm (maximum) | Operational power range | 1000 W |  | x
Accuracy | Workload accuracy below 100 W | Accuracy of the ergometer below 100 Watt load | 3 W |  | x
| Workload accuracy from 100 to 500 W | Accuracy of the ergometer between 100 and 500 Watt | 3 % |  | x
| Workload accuracy from 500 to 1000 W | Accuracy of the ergometer between 500 and 1000 Watt | 5 % |  | x
Comfort | Q-factor | Q-factor | 180 mm |  | x
| Minimum leg length user | Leg length to fit to pedals | 645 mm | 25.4 inch | x
| Minimum leg length user (incl. adjustable pedals) | Leg length to fit to pedals when adjustable cranks are used (measured inseam) | 602 mm | 23.7 inch | x
| Allowed user weight | Maximum patient weight | 180 kg | 396.8 lbs | x
| Handlebar adjustment angle | Angle over which the handlebar can rotate | 360 ° |  | x
| Adjustability range seat | Seat adjustment range | 300 mm | 11.8 inch | x
User Interface | English user interface | User interface available in English language | yes |  | x
| Norwegian user interface | User interface available in Norwegian language | yes |  | x
| Czech user interface | User interface available in Czech language | yes |  | x
| Danish user interface | User interface available in Danishlanguage | yes |  | x
| Dutch user interface | User interface available in Dutch language | yes |  | x
| Finnish user interface | User interface available in Finnish language | yes |  | x
| French user interface | User interface available in French language | yes |  | x
| German user interface | User interface available in German language | yes |  | x
| Italian user interface | User interface available in Italian language | yes |  | x
| Japanese user interface | User interface available in language | yes |  | x
| Korean user interface | User interface available in Korean language | yes |  | x
| Polish user interface | User interface available in Polish language | yes |  | x
| Portugese user interface | User interface available in Portugese language | yes |  | x
| Russian user interface | User interface available in Russian language | yes |  | x
| Spanish user interface | User interface available in Spanish language | yes |  | x
| Turkish user interface | User interface available in Turkish language | yes |  | x
| Ukrainian user interface | User interface available in Ukrainian language | yes |  | x
| Readout RPM | Readout the actual rpm on control unit | yes |  | x
| Readout Time | Readout time on control unit | yes |  | x
| Readout Power | Readout the actual power on control unit | yes |  | x
| Set Resistance | Possibility to define a specific resistance for each test subject | yes |  | x
| Terminal operation mode | Operate with an external device | yes |  | x
| Screen size (diagonal) | Diagonal size of the display | 8.9 cm | 3.5 inch | x
| Touchscreen | Operation by touch screen | yes |  | x
Connectivity | Lode 38K4 interface protocol | Supports communication with Lode 38K4 protocol | yes |  | x
| Lode interface protocol | Supports communication according to the Lode ergometer protocol | yes |  | x
| Lode WLP interface protocol | Supports communication with Lode WLP protocol | yes |  | x
| Ergoline P10 interface protocol | Supports communication with Ergoline P10 protocol | yes |  | x
| Ergoline P4 interface protocol | Supports communication with Ergoline P4 protocol | yes |  | x
| Schiller interface protocol | Schiller-käyttöliittymäprotokolla | yes |  | x
| Bosch EKG 506 DS interface protocol | Supports communication with Bosch EKG 506 DS Protokoll | yes |  | x
| USB connector | Connection via USB possible | yes |  | x
| RS232 in connector | Connection via RS232 | yes |  | x
Dimensions | Screen resolution | Screen resolution | 320 x 240 pixels |  | x
| Product length (cm) | Length of product in cm | 105 cm | 41.3 inch | x
| Product width (cm) | Width of the product in cm | 46 cm | 18.1 inch | x
| Product height | Height of product | 114 cm | 44.9 inch | x
| Product weight | The weight of the product | 65 kg | 143.3 lbs | x
Power requirements | V AC | Voltage in Volt | 100 - 240 V |  | x
| Phases | Phases | 1 |  | x
| Frequency | Frequency in Herz | 50/60 Hz |  | x
| Power consumption | Power consumption in Watt | 160 W |  | x
| Power cord length | Length of powercord including plugs | 250 cm | 98.4 inch | x
| Power cord IEC 60320 C13 with CEE 7/7 plug | European power cord and connector standard supplied with product | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
Standards & Safety | IEC 60601-1:2012 | The product is IEC 60601-1 edition 3.1 compliant | yes |  | x
| ISO 13485:2016 compliant | Lode fulfils ISO 13485:2016 requirements | yes |  | x
| ISO 9001:2015 compliant | Lode fulfills ISO 9001: 2015 requirements | yes |  | x
Certification | CE class Im according to MDD93/42/EEC | CE certified | yes |  | x
| CE class of product with optional SpO2 | CE class of product with optional SpO2 | IIa |  | x
| CE class of product with optional BPM | CE class of product with optional Blood Pressure Measurement | IIa |  | x
| CB according to IECEE CB | CB certified | yes |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 35 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 106 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 80 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | 60 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | -20 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 95 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 10 % |  | x

## Required options
Name |  | Partnumber | Required
Control Unit with 7" touch screen for ergometer | Multifunctionality | 945834 | x
Table foot for Control Unit with Touchscreen | Control from any desired place | 945818 | x
Cable for separate Control Unit Treadmills | Multifunctionality | 945813 | x
Colour Display 3.5" - 2nd screen | Multifunctionality | 945819 | x
Add program function to 7" touch screen for ergometer | Easily programmable | U945835 | x
Bluetooth Smart heart rate | Available in Q2 2019 | 945833 | x
Control Unit with touch screen 7" - ordered additionally | Multifunctionality | P945834 | x
Programmable Control Unit with 7" Touchscreen for ergometer | Programmable | 945835 | x
Blood Pressure with ECG trigger for bicycle ergometer | with ECG trigger | 945828 | x
Blood Pressure Custom pack 01 | For GE Case - GE Cardiosoft ECG CAM 14 - Case 8000 - Sensormedics Max 1 - Mac 5000 | 945881 | x
Blood Pressure Custom pack 02 | For Burdick Quest - Cambridge Heart CH2000 (V1.5 to 2.0.3) - Cosmed Quark T/C12x - Marquette Case 12, 15, 16 / Centra - Mortara X-Scribe II - Norav - Philips Stress Vue | 945882 | x
BPM Custom pack 03 | For Marquette MAC VU Digital ECG | 945883 | x
BPM Custom pack 04 | For Marquette Hellige Cardiosys | 945884 | x
BPM Custom pack 05 | For Quinton Digital ECG - Nasiff Associates Cardio-Card | 945885 | x
BPM Custom pack 06 | For Welch Allyn CardioPerfect Workstation | 945886 | x
BPM Custom pack 07 | For Reynolds CardioDirect 12-S | 945887 | x
BPM Custom pack 08 | For HR triggering by belt | 945888 | x
BPM Custom pack 09 | For Schiller CS200 | 945889 | x
BPM Custom pack 10 | For Customed MC3000 | 945890 | x
Pediatric Cuff Complete 1.5 m | Easy to apply | 928872 | x
Adult Cuff - cuff only - | Easy to apply | 928987 | x
Microphone set for Cuff (1.5m) | Easy to apply | 928860 | x
Cuff Extra Large - cuff only - | Easy to apply | 928989 | x
Pediatric Cuff - Cuff only - | Easy to apply | 928988 | x
Cuff Complete 1.5 m | Easy to apply | 928870 | x
Cuff Extra Large Complete 1.5 m | Easy to apply | 928874 | x
ECG Acquisition trigger set with trigger | High Quality ECG Acquisition | 950824 | x
Electric adjustable saddle height | Easy and accurate positioning | 960810 | x
0-Watt start-up system | Lowest possible startup power | 960805 | x
Saddle for children | Versatile ergometry | 401066 | x
Saddle for children - ordered additionally | Versatile ergometry | P401066 | x
Ambient sensor pack | Check environmental conditions during test | 945827 | x
Adjustable cranks | Optimal force application | 928804 | x
Shortened saddle shaft | Increase flexibility for smaller people | 960806 | x
Easy saddle exchange option - Corival | Fast change of saddle to suit all users | 960807 | x
Saddle extra large | Versatile ergometry | 401084 | x
Pedal shoes (pair) | Extra stability during cycling | 917803 | x
Pedal shoes pediatric (pair) | Pedal shoes for childen | 917833 | x
Pedal shoes extra large (pair) | For large feet sizes | 917834 | x
Transportwheel for Corival | Easy transportation indoors | 960801 | x
RS232 cable | Easy connection | 930911 | x
USB to Serial converter | Easy connection | 226012 | x
Arm support | Arterial line possible | 906814 | x
SpO2 for control unit with touch panel (bicycle) | Saturation and heart rate | 945823 | x
Adult Finger clip, 1 meter cable | Easy to apply and clean | 928946 | x
Adult Finger clip, 2 meter cable | Easy to apply and clean | 928947 | x
Adult Finger clip, 3 meter cable | Easy to apply and clean | 928948 | x
Earclip sensor, 1 meter cable | Easy to apply and clean | 928951 | x
Reflectance sensor, 1 meter cable | Easy to apply and clean | 928952 | x
Reflectance sensor holder, 10 caps & 20 adhesive stickers | Easy to apply and clean | 928953 | x
Adult flex sensor with 25 flexiwrap, 1 meter cable | Easy to apply and clean | 928954 | x
Adult flex sensor with 25 flexiwrap, 3 meter cable | Easy to apply and clean | 928955 | x
Sensor Adapter Cable, 18 cm length, 8 to 9 pin connector | Versatile connectivity | 928961 | x
Reusable Soft Sensor Medium, 1 meter cable | New technology | 928967 | x
Reusable Soft Sensor Medium, 3 meter cable | New technology | 928969 | x
SpO2 for control unit with touch panel - ordered afterwards | Ordered afterwards | P945823 | x

## Spareparts
Name | Partnumber
3,5" Control Unit front panel | 945110
7" Control Unit front panel | F945105
7" Programmable Control Unit front panel | F945105P
AC Cable - C13 2,5 Metres | 212010
Assembly seat adjustment | 960030
Click fastener | 871826
Cover Left | 960060
Cover Right | 960061
Cover top side | 960063
Crank set Corival family | 401175
Drive Chain | 960080
Handlebar cosmetic cover | 960056
Handlebar Foam Corival cpet/rehab | 960098
High Performance Ergometer Oil | F928601
Main switch and cable | 960703
Pedals with straps | 4011012
Plastic foot | 871899
Power cable ECM | 960704
Power supply | 960040
Saddle | 4010392