## Specifications
Product | Angio cpet - with static wall fixation
Partnumber | 967902
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Workload range fixed torque | Workload range in fixed torque mode | 0,1 - 70 Nm |  | x
| Minimum load | The minimum load the ergometer can provide | 7 W |  | x
| Maximum peak load | The maximum load the ergometer can provide for a short period of time | 1000 W |  | x
| Minimum load increments | The smallest steps with which the load can be added | 1 W |  | x
| Maximum continuous load | The maximum power the ergometer can deliver continuous | 750 W |  | x
| Hyperbolic workload control | Control of workload in hyperbolic way | yes |  | x
| Linear workload control | Control ergometer in a linear way | yes |  | x
| Fixed torque workload control | Control ergometer in a fixed torque way | yes |  | x
| Maximum rpm independent constant load | Rpm as to which constant load can be applied | 150 rpm |  | x
| Minimum rpm independent constant load | Rpm as from which constant load can be applied | 30 rpm |  | x
| Optional heart rate controlled workload | Control the device optionally through a set heartrate | yes |  | x
| Electromagnetic "eddy current" braking system | System that very accurately applies a brake to the ergometer | yes |  | x
| Dynamic calibration | Continuous calibration during application of workload | yes |  | x
Accuracy | Workload accuracy from 7 to 100 W | Accuracy of the ergometer between 7 and 100 Watt | 3 W |  | x
| Workload accuracy from 100 to 500 W | Accuracy of the ergometer between 100 and 500 Watt | 3 % |  | x
| Workload accuracy from 500 to 1000 W | Accuracy of the ergometer between 500 and 1000 Watt | 5 % |  | x
User Interface | Readout Distance | Readout the distance on control unit | yes |  | x
| Readout RPM | Readout the actual rpm on control unit | yes |  | x
| Readout Heartrate | Readout the actual heartarte on control unit | yes |  | x
| Readout target HR | Readout target heartarte on control unit (only possible with HR option) | yes |  | x
| Readout Energy | Readout the energy on control unit | yes |  | x
| Readout Torque | Readout torque on control unit | yes |  | x
| Readout Time | Readout time on control unit | yes |  | x
| Readout Power | Readout the actual power on control unit | yes |  | x
| Set Display | Choosing parameters to be shown on the display during the stress test | yes |  | x
| Set Resistance | Possibility to define a specific resistance for each test subject | yes |  | x
| Set P-Slope | Possibility to set the load regulation resistance | yes |  | x
| Set Mode | Possibility to define parameters to be adjusted during the stress test | yes |  | x
| Manual operation mode | Operate manually | yes |  | x
| Preset protocol operation mode | Operate via preset protocols | yes |  | x
| Terminal operation mode | Operate with an external device | yes |  | x
| External control unit | A control unit is connected with a cable to the ergometer | yes |  | x
| Selfdesigned protocol operation mode | Operate via selfdesigned protocols | yes |  | x
Dimensions | Product length (cm) | Length of product in cm | 54 cm | 21.3 inch | x
| Product width (cm) | Width of the product in cm | 68 cm | 26.8 inch | x
| Product height | Height of product | 73 cm | 28.7 inch | x
| Product weight | The weight of the product | 55 kg | 121.3 lbs | x
Power requirements | V AC | Voltage in Volt | 100 - 240 V |  | x
| Phases | Phases | 1 |  | x
| Frequency | Frequency in Herz | 50/60 Hz |  | x
| Power consumption | Power consumption in Watt | 160 W |  | x
| Power cord IEC 60320 C13 with CEE 7/7 plug | European power cord and connector standard supplied with product | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
Standards & Safety | IEC 60601-1:2012 | The product is IEC 60601-1 edition 3.1 compliant | yes |  | x
| ISO 13485:2016 compliant | Lode fulfils ISO 13485:2016 requirements | yes |  | x
| ISO 9001:2015 compliant | Lode fulfills ISO 9001: 2015 requirements | yes |  | x
Certification | CE class Im according to MDD93/42/EEC | CE certified | yes |  | x
| CE class of product with optional SpO2 | CE class of product with optional SpO2 | IIa |  | x
| CE class of product with optional BPM | CE class of product with optional Blood Pressure Measurement | IIa |  | x
| CB according to IECEE CB | CB certified | yes |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 35 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 106 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 80 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | 60 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | -20 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 95 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 10 % |  | x

## Required options
Name |  | Partnumber | Required
Control Unit with 7" touch screen for ergometer | Multifunctionality | 945834 | x
Table foot for Control Unit with Touchscreen | Control from any desired place | 945818 | x
Cable for separate Control Unit Treadmills | Multifunctionality | 945813 | x
Colour Display 3.5" - 2nd screen | Multifunctionality | 945819 | x
Add program function to 7" touch screen for ergometer | Easily programmable | U945835 | x
Bluetooth Smart heart rate | Available in Q2 2019 | 945833 | x
Programmable Control Unit with 7" Touchscreen for ergometer | Programmable | 945835 | x
SpO2 for control unit with touch panel (bicycle) | Saturation and heart rate | 945823 | x
Adult Finger clip, 1 meter cable | Easy to apply and clean | 928946 | x
Adult Finger clip, 2 meter cable | Easy to apply and clean | 928947 | x
Adult Finger clip, 3 meter cable | Easy to apply and clean | 928948 | x
Earclip sensor, 1 meter cable | Easy to apply and clean | 928951 | x
Reflectance sensor, 1 meter cable | Easy to apply and clean | 928952 | x
Reflectance sensor holder, 10 caps & 20 adhesive stickers | Easy to apply and clean | 928953 | x
Adult flex sensor with 25 flexiwrap, 1 meter cable | Easy to apply and clean | 928954 | x
Adult flex sensor with 25 flexiwrap, 3 meter cable | Easy to apply and clean | 928955 | x
Sensor Adapter Cable, 18 cm length, 8 to 9 pin connector | Versatile connectivity | 928961 | x
Reusable Soft Sensor Medium, 1 meter cable | New technology | 928967 | x
Reusable Soft Sensor Medium, 3 meter cable | New technology | 928969 | x
Adjustable cranks | Optimal force application | 928804 | x
USB to Serial converter | Easy connection | 226012 | x