## Specifications
Product | Brachumera sport
Partnumber | 925901
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Minimum load | The minimum load the ergometer can provide | 7 W |  | x
| Maximum peak load | The maximum load the ergometer can provide for a short period of time | 2500 W |  | x
| Isokinetic workload control | Control ergometer in isokinetic way | yes |  | x
| Minimum load increments | The smallest steps with which the load can be added | 1 W |  | x
| Maximum continuous load | The maximum power the ergometer can deliver continuous | 1500 W |  | x
| Hyperbolic workload control | Control of workload in hyperbolic way | yes |  | x
| Linear workload control | Control ergometer in a linear way | yes |  | x
| Fixed torque workload control | Control ergometer in a fixed torque way | yes |  | x
| Maximum rpm independent constant load | Rpm as to which constant load can be applied | 150 rpm |  | x
| Minimum rpm independent constant load | Rpm as from which constant load can be applied | 30 rpm |  | x
| Optional heart rate controlled workload | Control the device optionally through a set heartrate | yes |  | x
| Electromagnetic "eddy current" braking system | System that very accurately applies a brake to the ergometer | yes |  | x
| Dynamic calibration | Continuous calibration during application of workload | yes |  | x
Accuracy | Workload accuracy below 100 W | Accuracy of the ergometer below 100 Watt load | 3 W |  | x
| Workload accuracy from 100 to 1500 W | Accuracy of the ergometer between 100 and 1500 Watt | 3 % |  | x
| Workload accuracy over 1500 W | Accuracy of the ergometer over 1500 Watt | 5 % |  | x
User Interface | Readout Distance | Readout the distance on control unit | yes |  | x
| Readout RPM | Readout the actual rpm on control unit | yes |  | x
| Readout Heartrate | Readout the actual heartarte on control unit | yes |  | x
| Readout target HR | Readout target heartarte on control unit (only possible with HR option) | yes |  | x
| Readout Energy | Readout the energy on control unit | yes |  | x
| Readout Torque | Readout torque on control unit | yes |  | x
| Readout Time | Readout time on control unit | yes |  | x
| Readout Power | Readout the actual power on control unit | yes |  | x
| Set Display | Choosing parameters to be shown on the display during the stress test | yes |  | x
| Set Resistance | Possibility to define a specific resistance for each test subject | yes |  | x
| Set P-Slope | Possibility to set the load regulation resistance | yes |  | x
| Set Mode | Possibility to define parameters to be adjusted during the stress test | yes |  | x
| Manual operation mode | Operate manually | yes |  | x
| Preset protocol operation mode | Operate via preset protocols | yes |  | x
| Analog operation mode | Operate analog | yes |  | x
| Terminal operation mode | Operate with an external device | yes |  | x
| Selfdesigned protocol operation mode | Operate via selfdesigned protocols | yes |  | x
Connectivity | Analog connector | Read-out of ergometer load in an analog way | yes |  | x
| RS232 in connector | Connection via RS232 | yes |  | x
Dimensions | Product length (cm) | Length of product in cm | 114 cm | 44.9 inch | x
| Product width (cm) | Width of the product in cm | 59 cm | 23.2 inch | x
| Product height | Height of product | 51 cm | 20.1 inch | x
| Product weight | The weight of the product | 65 kg | 143.3 lbs | x
Power requirements | Power cord length | Length of powercord including plugs | 250 cm | 98.4 inch | x
| Power cord IEC 60320 C13 with CEE 7/7 plug | European power cord and connector standard supplied with product | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
| 115 V AC 50/60 Hz (130 VA) | Operational power | yes |  | x
| 230 V AC 50/60 Hz (130 VA) | Operational power | yes |  | x
Standards & Safety | IEC 60601-1:2005 | The product is IEC 60601-1 3rd edition compliant | yes |  | x
| ISO 13485:2016 compliant | Lode fulfils ISO 13485:2016 requirements | yes |  | x
| ISO 9001:2015 compliant | Lode fulfills ISO 9001: 2015 requirements | yes |  | x
Certification | CE class Im according to MDD93/42/EEC | CE certified | yes |  | x
| CE class of product with optional SpO2 | CE class of product with optional SpO2 | IIa |  | x
| CE class of product with optional BPM | CE class of product with optional Blood Pressure Measurement | IIa |  | x
| CB according to IECEE CB | CB certified | yes |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 40 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 106 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 70 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | 70 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | -25 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 95 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 10 % |  | x

## Required options
Name |  | Partnumber | Required
Programmable Control Unit | Programming protocols in advance | 928811 | x
Programmable Control Unit ordered afterwards | ordered afterwards | P928811 | x
Programmable Control Unit with SpO2 & Heart rate | Measurement of oxygen saturation during stress testing. | 928841 | x
Adult Finger clip, 1 meter cable | Easy to apply and clean | 928946 | x
Adult Finger clip, 2 meter cable | Easy to apply and clean | 928947 | x
Adult Finger clip, 3 meter cable | Easy to apply and clean | 928948 | x
Pediatric Finger clip, 1 meter cable | Easy to apply and clean | 928949 | x
Pediatric Finger clip, 3 meter cable | Easy to apply and clean | 928950 | x
Earclip sensor, 1 meter cable | Easy to apply and clean | 928951 | x
Adult flex sensor with 25 flexiwrap, 1 meter cable | Easy to apply and clean | 928954 | x
Adult flex sensor with 25 flexiwrap, 3 meter cable | Easy to apply and clean | 928955 | x
Flexi Wrap - Adult, package of 25 | Easy to apply and clean | 928958 | x
Flexi Wrap - Infant, package of 25 | Easy to apply and clean | 928957 | x
Sensor Adapter Cable, 18 cm length, 8 to 9 pin connector | Versatile connectivity | 928961 | x
Right Angle Connector Cable, 15 cm length | Versatile connectivity | 928962 | x
Right Angle Connector Cable, 3 m length | Versatile connectivity | 928963 | x
Patient Extension Cable, 1 m length | Versatile connectivity | 928964 | x
Patient Extension Cable, 3 m length | Versatile connectivity | 928965 | x
Reusable Soft Sensor Medium, 1 meter cable | New technology | 928967 | x
Reusable Soft Sensor Medium, 3 meter cable | New technology | 928969 | x
Programmable Control Unit with SpO2 & Heart rate ordered afterwards | ordered afterwards | P928841 | x
Heart rate | Heart rate controlled cycling | 928826 | x
Heart rate (ordered afterwards) | ordered afterwards | P928826 | x
0-Watt start-up system | Lowest possible startup power | 925805 | x
Adjustable sports cranks incl. pediatric range | Optimal force application | 925808 | x
Adjustable sports cranks incl. pediatric range - ordered afterwards | ordered afterwards | P925808 | x
Adjustable wall fixation for Brachumera sport | Versatile application of Brachumera Sport | 925830 | x
Electric adjustable chair for arm ergometer | Comfortable seating position in front of the ergometer | 917813 | x
Floor fixation set for adjustable chair for arm ergometer | Extra stability because of floor fixation | 917815 | x
Stand for adjustable wall fixation Brachumera sport | Versatile application of Brachumera sport | 925840 | x
USB to Serial converter | Easy connection | 226012 | x
RS232 cable | Easy connection | 930911 | x
Bluetooth Smart heart rate | Available in Q2 2019 | 945833 | x