## Specifications
Product | Esseda Wheelchair ergometer
Partnumber | 988900
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Maximum speed | Maximum speed of the treadmill | 45 km/h | 28.0 mph | x
| Minimum operational speed | Minimum speed of the treadmill | 0.1 km/h | 0.1 mph | x
| Speed adjustment steps | The steps you can adjust the speed with | 0.1 km/h | 0.1 mph | x
| Maximum reverse speed | Maximum speed of the treadmill in reverse | 45 km/h | 28.0 mph | x
| Rotation direction clockwise | Looking at the outgoing axis | yes |  | x
| Rotation direction anti-clockwise | Looking at the outgoing axis | yes |  | x
Accuracy | Speed readout accuracy | The speed readout accuracy of the treadmill | 0.1 km/h | 0.1 mph | x
Comfort | Allowed user weight | Maximum patient weight | 225 kg | 496.0 lbs | x
Connectivity | High Speed data input in LEM | High frequency datastream from measurement cell direct into Lode Ergometry Manager | yes |  | x
| USB connector | Connection via USB possible | yes |  | x
Dimensions | Product height (mm) | Height of product in mm | 220 mm |  | x
| Product length (mm) | Length of the product in mm | 3311 mm |  | x
| Product width (mm) | Width of the product in mm | 1394 mm |  | x
| Product weight | The weight of the product | 200 kg | 440.9 lbs | x
Power requirements | Power cord length | Length of powercord including plugs | 300 cm | 118.1 inch | x
| Power cord IEC 60320 C19 with CEE 7/7 plug | European power cord and connector standard supplied with product | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
| 230 V AC 50/60 Hz | Operational power | yes |  | x
Standards & Safety | Integrated emergency stop | Emergency stop fully integrated | yes |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 35 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 106 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 80 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | -20 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | 60 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 10 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 95 % |  | x
Included parts | PC software included | Software needed for operation included | yes |  | x
| All cables included | Cables for installing the complete setup are included | yes |  | x

## Required options
Name |  | Partnumber | Required
LEM 10 - Module Export | Export your data to MS Excel | 955920 | x
LEM 10 - Module Architect | Creating protocols with customer specific formulas | 955921 | x