## Specifications
Product | Katana sport
Partnumber | 934900
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Maximum speed | Maximum speed of the treadmill | 30 km/h | 18.6 mph | x
| Minimum operational speed | Minimum speed of the treadmill | 0.1 km/h | 0.1 mph | x
| Speed adjustment steps | The steps you can adjust the speed with | 0.1 km/h | 0.1 mph | x
| Positive elevation | Maximum positive angle as a percentage | 25 % |  | x
| Elevation adjustment steps | The size of the steps you can adjust the elevation | 0.5 % |  | x
| Optional negative elevation | Mimimum negative angle percentage that can be reached with the optional accessory | -25 % |  | x
Accuracy | Elevation readout accuracy | The elevation readout accuracy of the treadmill | 0.3 % |  | x
| Speed readout accuracy | The speed readout accuracy of the treadmill | 0.5 km/h | 0.3 mph | x
Comfort | Allowed user weight | Maximum patient weight | 225 kg | 496.0 lbs | x
Connectivity | Lode ERM interface protocol | Supports communication according to the Lode ERM treadmill protocol | yes |  | x
| "HP Cosmos" interface protocol | Supports communication according to the "HP Cosmos" protocol | yes |  | x
| "Woodway" interface protocol | Supports communication according to the "Woodway" protocol | yes |  | x
| "Trackmaster" interface protocol | Supports communication according to the "Trackmaster" protocol | yes |  | x
| "GE" interface protocol | Supports communication according to the "GE" protocol | yes |  | x
| USB connector | Connection via USB possible | yes |  | x
Dimensions | Screen resolution | Screen resolution | 800 x 400 pixels |  | x
| Walking surface length | Length of the walking surface | 190 cm | 74.8 inch | x
| Walking surface width | Width of the walking surface | 70 cm | 27.6 inch | x
| Step up height | Height to step up the treadmill | 29 cm | 11.4 inch | x
| Product length (cm) | Length of product in cm | 250 cm | 98.4 inch | x
| Product width (cm) | Width of the product in cm | 106 cm | 41.7 inch | x
| Product height | Height of product | 145 cm | 57.1 inch | x
| Product weight | The weight of the product | 380 kg | 837.8 lbs | x
Power requirements | Power cord 3L+N+PE | Permanently installed device | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
| Maximum rated power input | Power input | 5000 VA |  | x
| 3 x 400 V | Needed power connection | yes |  | x
| Maximum motor power | The maximum power the motor will consume | 3.1 kW |  | x
Standards & Safety | IEC 60601-1:2005 | The product is IEC 60601-1 3rd edition compliant | yes |  | x
| ISO 13485:2016 compliant | Lode fulfils ISO 13485:2016 requirements | yes |  | x
| ISO 9001:2015 compliant | Lode fulfills ISO 9001: 2015 requirements | yes |  | x
| Standard emergency lanyard | An emergency stop with a magnetic lanyard is present | yes |  | x
| Step up plate emergency stop | Emergency stop when touching the step up plate | yes |  | x
Certification | CE class Im according to MDD93/42/EEC | CE certified | yes |  | x
| CE class of product with optional SpO2 | CE class of product with optional SpO2 | IIa |  | x
| CE class of product with optional BPM | CE class of product with optional Blood Pressure Measurement | IIa |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 35 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 10 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 70 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | 70 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | -25 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 95 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 10 % |  | x

## Required options
Name |  | Partnumber | Required
Remote Control for treadmill | Ease of use | 945831 | x
Add program function to 7" touch screen for Treadmill | Programmable | U945815 | x
Remote Control for treadmill (ordered afterwards) | Ease of use | P945831 | x
Table foot for Control Unit with Touchscreen | Control from any desired place | 945818 | x
Cable for separate Control Unit Treadmills | Multifunctionality | 945813 | x
Colour Display 3.5" - 2nd screen | Multifunctionality | 945819 | x
SpO2 for control unit with touch panel (treadmill) | Oxygen saturation | 945822 | x
Adult Finger clip, 1 meter cable | Easy to apply and clean | 928946 | x
Adult Finger clip, 2 meter cable | Easy to apply and clean | 928947 | x
Adult Finger clip, 3 meter cable | Easy to apply and clean | 928948 | x
Earclip sensor, 1 meter cable | Easy to apply and clean | 928951 | x
Reflectance sensor, 1 meter cable | Easy to apply and clean | 928952 | x
Reflectance sensor holder, 10 caps & 20 adhesive stickers | Easy to apply and clean | 928953 | x
Adult flex sensor with 25 flexiwrap, 1 meter cable | Easy to apply and clean | 928954 | x
Adult flex sensor with 25 flexiwrap, 3 meter cable | Easy to apply and clean | 928955 | x
Sensor Adapter Cable, 18 cm length, 8 to 9 pin connector | Versatile connectivity | 928961 | x
Reusable Soft Sensor Medium, 1 meter cable | New technology | 928967 | x
Reusable Soft Sensor Medium, 3 meter cable | New technology | 928969 | x
SpO2 for control unit with touch panel - ordered afterwards | Ordered afterwards | P945822 | x
Blood Pressure Measurement with ECG trigger for treadmills | with ECG trigger | 945824 | x
Blood Pressure Custom pack 01 | For GE Case - GE Cardiosoft ECG CAM 14 - Case 8000 - Sensormedics Max 1 - Mac 5000 | 945881 | x
Blood Pressure Custom pack 02 | For Burdick Quest - Cambridge Heart CH2000 (V1.5 to 2.0.3) - Cosmed Quark T/C12x - Marquette Case 12, 15, 16 / Centra - Mortara X-Scribe II - Norav - Philips Stress Vue | 945882 | x
BPM Custom pack 03 | For Marquette MAC VU Digital ECG | 945883 | x
BPM Custom pack 04 | For Marquette Hellige Cardiosys | 945884 | x
BPM Custom pack 05 | For Quinton Digital ECG - Nasiff Associates Cardio-Card | 945885 | x
BPM Custom pack 06 | For Welch Allyn CardioPerfect Workstation | 945886 | x
BPM Custom pack 07 | For Reynolds CardioDirect 12-S | 945887 | x
BPM Custom pack 09 | For Schiller CS200 | 945889 | x
Cuff Complete Adult - small (18-27 cm) | easy to apply | 945870 | x
Cuff Complete Adult (25-35 cm) | easy to apply | 945871 | x
Cuff Complete Adult - plus (27-40 cm) | easy to apply | 945872 | x
Cuff Complete Adult - large (32-44 cm) | easy to apply | 945873 | x
Rear Elevation -25% electrical Katana sport | Possibility of downhill walking | 934808 | x
Emergency Stop Button | Ultimate safety | 945804 | x
Autospeed | Flexibility in exercising | 945840 | x
Adjustable Acceleration & Deceleration for Katana & Ultra | Flexibility in exercising | 945847 | x
Speed 40 km/h | Extra high speed | 934840 | x
Reverse walking for Katana | walking backwards | 934842 | x
Ambient sensor pack | Check environmental conditions during test | 945826 | x
Safety Belt and Fall stop for Katana sport | Extra safety | 934821 | x
Harness Fallstop Small | Suitable for smaller people | 458102 | x
Harness Fallstop XL | Suitable for larger people | 458105 | x
Harness Fallstop XS | Suitable for children | 458101 | x
Harness Fallstop Large | Suitable for larger people | 458104 | x
Harness Fallstop Medium | Standard supplied with Fallstop | 458103 | x
Extension for emergency lanyard | Original Accessory | 945931 | x
Handrail, removable - for Katana sport | Convenient removal of handrail for many purposes | 934811 | x
Universal treadmill Arm Support | Comfort for both test subject and test operator | 945805 | x
USB to Serial converter | Easy connection | 226012 | x
RS232 cable | Easy connection | 930911 | x
Network module connection cable | Easy connection | 930930 | x
Bluetooth Smart heart rate | Available in Q2 2019 | 945833 | x