## Specifications
Product | Valiant Ultra 250
Partnumber | 938925
| Specification item | Description | Requirement | Imperial units | Fullfills requirement
Workload | Maximum speed | Maximum speed of the treadmill | 40 km/h | 24.9 mph | x
| Minimum operational speed | Minimum speed of the treadmill | 1 km/h | 0.6 mph | x
| Speed adjustment steps | The steps you can adjust the speed with | 0.1 km/h | 0.1 mph | x
| Positive elevation | Maximum positive angle as a percentage | 25 % |  | x
| Negative elevation | Maximum negative angle percentage that can be reached | -25 % |  | x
Comfort | Allowed user weight | Maximum patient weight | 225 kg | 496.0 lbs | x
User Interface | English user interface | User interface available in English language | yes |  | x
| Norwegian user interface | User interface available in Norwegian language | yes |  | x
| Czech user interface | User interface available in Czech language | yes |  | x
| Danish user interface | User interface available in Danishlanguage | yes |  | x
| Dutch user interface | User interface available in Dutch language | yes |  | x
| Finnish user interface | User interface available in Finnish language | yes |  | x
| French user interface | User interface available in French language | yes |  | x
| German user interface | User interface available in German language | yes |  | x
| Italian user interface | User interface available in Italian language | yes |  | x
| Japanese user interface | User interface available in language | yes |  | x
| Korean user interface | User interface available in Korean language | yes |  | x
| Polish user interface | User interface available in Polish language | yes |  | x
| Portugese user interface | User interface available in Portugese language | yes |  | x
| Russian user interface | User interface available in Russian language | yes |  | x
| Spanish user interface | User interface available in Spanish language | yes |  | x
| Turkish user interface | User interface available in Turkish language | yes |  | x
| Ukrainian user interface | User interface available in Ukrainian language | yes |  | x
| Readout Angle | Readout actual elevation on control unit (treadmill only) | yes |  | x
| Readout Target Angle | Readout the target elevation on control unit (treadmill only) | yes |  | x
| Readout Distance | Readout the distance on control unit | yes |  | x
| Readout Speed | Readout the actual speed on control unit | yes |  | x
| Readout Energy | Readout the energy on control unit | yes |  | x
| Readout Time | Readout time on control unit | yes |  | x
| Set Display | Choosing parameters to be shown on the display during the stress test | yes |  | x
| Customized Screenviews | Customized Screenviews | yes |  | x
| Manual operation mode | Operate manually | yes |  | x
| Quick start operation mode | Easy quick start possible | yes |  | x
| Touchscreen | Operation by touch screen | yes |  | x
Connectivity | Lode 38K4 interface protocol | Supports communication with Lode 38K4 protocol | yes |  | x
| Lode ERM interface protocol | Supports communication according to the Lode ERM treadmill protocol | yes |  | x
| "HP Cosmos" interface protocol | Supports communication according to the "HP Cosmos" protocol | yes |  | x
| "Woodway" interface protocol | Supports communication according to the "Woodway" protocol | yes |  | x
| "Trackmaster" interface protocol | Supports communication according to the "Trackmaster" protocol | yes |  | x
| "GE" interface protocol | Supports communication according to the "GE" protocol | yes |  | x
| USB connector | Connection via USB possible | yes |  | x
| RS232 in connector | Connection via RS232 | yes |  | x
Dimensions | Screen resolution | Screen resolution | 800 x 400 pixels |  | x
| Walking surface length | Length of the walking surface | 250 cm | 98.4 inch | x
| Walking surface width | Width of the walking surface | 150 cm | 59.1 inch | x
| Step up height | Height to step up the treadmill | 35 cm | 13.8 inch | x
| Product length (cm) | Length of product in cm | 254 cm | 100.0 inch | x
| Product width (cm) | Width of the product in cm | 246 cm | 96.9 inch | x
| Product height | Height of product | 130 cm | 51.2 inch | x
| Product weight | The weight of the product | 850 kg | 1,873.9 lbs | x
Power requirements | Power cord 3L+N+PE | Permanently installed device | yes |  | x
| Power cord NEMA | NEMA power cord and connector | no |  | x
| Maximum rated power input | Power input | 8000 VA |  | x
| 3 x 400 V | Needed power connection | yes |  | x
| Maximum motor power | The maximum power the motor will consume | 8 kW |  | x
Standards & Safety | ISO 9001:2015 compliant | Lode fulfills ISO 9001: 2015 requirements | yes |  | x
| Integrated emergency stop | Emergency stop fully integrated | yes |  | x
| Standard emergency lanyard | An emergency stop with a magnetic lanyard is present | yes |  | x
Environmental conditions | Maximum operational temperature | Maximum temperature at which the device will work within specification | 40 °C |  | x
| Minimum operational temperature | Minimum temperature at which the device will work within specification | 14 °C |  | x
| Maximum operational air pressure | Maximum air pressure as from which the product will be operating within specification | 106 kPa |  | x
| Minimum operational air pressure | Minimum airpressure as from which the product will be operating within specification | 70 kPa |  | x
| Maximum operational non-condensing humidity | Maximum non-condensing humidity at which the product will operate within specification | 90 % |  | x
| Minimum operational non-condensing humidity | Minimum non-condensing humidity at which the product will operate within specification | 30 % |  | x
| Maximum storage & transport temperature | Maximum temperature the product may be stored or transported at | 70 °C |  | x
| Minimum storage & transport temperature | Minimum temperature the product may be stored and transported at | -25 °C |  | x
| Maximum storage & transport air pressure | Maximum air pressure the product may be stored and transported at | 106 kPa |  | x
| Minimum air pressure storage & transport | Minimum air pressure the product may be stored or transported at | 50 kPa |  | x
| Max. humidity storage & transport | Max. non-condensing humidity that the product can be stored or transported at | 95 % |  | x
| Min. humidity storage & transport | Minimum non-condensing humidity that the product can be stored or transport at | 10 % |  | x

## Required options
Name |  | Partnumber | Required
Remote Control for treadmill | Ease of use | 945831 | x
Colour Display 3.5" - 2nd screen | Multifunctionality | 945819 | x
SpO2 for control unit with touch panel (treadmill) | Oxygen saturation | 945822 | x
Adult Finger clip, 1 meter cable | Easy to apply and clean | 928946 | x
Adult Finger clip, 2 meter cable | Easy to apply and clean | 928947 | x
Adult Finger clip, 3 meter cable | Easy to apply and clean | 928948 | x
Earclip sensor, 1 meter cable | Easy to apply and clean | 928951 | x
Reflectance sensor, 1 meter cable | Easy to apply and clean | 928952 | x
Reflectance sensor holder, 10 caps & 20 adhesive stickers | Easy to apply and clean | 928953 | x
Adult flex sensor with 25 flexiwrap, 1 meter cable | Easy to apply and clean | 928954 | x
Adult flex sensor with 25 flexiwrap, 3 meter cable | Easy to apply and clean | 928955 | x
Sensor Adapter Cable, 18 cm length, 8 to 9 pin connector | Versatile connectivity | 928961 | x
Reusable Soft Sensor Medium, 1 meter cable | New technology | 928967 | x
Reusable Soft Sensor Medium, 3 meter cable | New technology | 928969 | x
SpO2 for control unit with touch panel - ordered afterwards | Ordered afterwards | P945822 | x
Blood Pressure Measurement with ECG trigger for treadmills | with ECG trigger | 945824 | x
Blood Pressure Custom pack 01 | For GE Case - GE Cardiosoft ECG CAM 14 - Case 8000 - Sensormedics Max 1 - Mac 5000 | 945881 | x
Blood Pressure Custom pack 02 | For Burdick Quest - Cambridge Heart CH2000 (V1.5 to 2.0.3) - Cosmed Quark T/C12x - Marquette Case 12, 15, 16 / Centra - Mortara X-Scribe II - Norav - Philips Stress Vue | 945882 | x
BPM Custom pack 03 | For Marquette MAC VU Digital ECG | 945883 | x
BPM Custom pack 04 | For Marquette Hellige Cardiosys | 945884 | x
BPM Custom pack 05 | For Quinton Digital ECG - Nasiff Associates Cardio-Card | 945885 | x
BPM Custom pack 06 | For Welch Allyn CardioPerfect Workstation | 945886 | x
BPM Custom pack 07 | For Reynolds CardioDirect 12-S | 945887 | x
BPM Custom pack 09 | For Schiller CS200 | 945889 | x
Cuff Complete Adult - small (18-27 cm) | easy to apply | 945870 | x
Cuff Complete Adult (25-35 cm) | easy to apply | 945871 | x
Cuff Complete Adult - plus (27-40 cm) | easy to apply | 945872 | x
Cuff Complete Adult - large (32-44 cm) | easy to apply | 945873 | x
Emergency Stop Button | Ultimate safety | 945804 | x
Autospeed | Flexibility in exercising | 945840 | x
Adjustable Acceleration & Deceleration for Katana & Ultra | Flexibility in exercising | 945847 | x
Ambient sensor pack | Check environmental conditions during test | 945826 | x
Bottle holder | Convenient beverage storage | 401200 | x
Extension for emergency lanyard | Original Accessory | 945931 | x
Table foot for Control Unit with Touchscreen | Control from any desired place | 945818 | x
Cable for separate Control Unit Treadmills | Multifunctionality | 945813 | x
Wheelchair support for Valiant Ultra 250 | Versatile wheelchair testing | 938870 | x
Entrance plate Valiant Ultra | Easy entrance | 938852 | x
Remote Control for treadmill (ordered afterwards) | Ease of use | P945831 | x
Bluetooth Smart heart rate | Available in Q2 2019 | 945833 | x